U.S. CUSTOMS AND BORDER PROTECTION

LABORATORY METHODS

CBPL METHOD 22-02

Method for the Determination of

Excess Pressure in Sparkling Wines

SAFETY STATEMENT

This CBPL Method cannot fully address safety issues that may arise from its use.  The analyst is responsible for assessing potential safety issues associated with a given method at its point of use.

Before using this method, the analyst will consider all general laboratory safety precautions.  In particular, the analyst will identify and implement suitable health and safety measures and will comply with all pertinent regulations.

METHOD UNCERTAINTY

The uncertainty of measurement for this method is specific to each laboratory.

0.		SCOPE AND FIELD OF 				APPLICATION

According to the Harmonized Tariff Schedule of the United States (HTSUS), Chapter 22, Subheading Note 1, "For the purposes of Subheading 2204.10 the expression sparkling wine means wine which, when kept at a temperature of 20 °C in closed containers, has an excess pressure of not less than 3 bars."  This method describes the use of commercially available aphrometers to measure excess pressure in wine bottles.

1.		REFERENCES

1.1		AOAC 960.20.  “Carbon Dioxide in Wines:  Manometric Method.”

1.2		AOAC 960.21.  “Carbon Dioxide in Wines:  Volumetric Method.”

##### 1.3		CBPL 22-05/AOAC 964.09.  “Carbon Dioxide in Wines:  Enzymatic Method.”

##### 1.4		CBPL 22-06/AOAC 988.07.  “Carbon Dioxide in Wines:  Titrimetric Method.”

1.5		Code of Federal Regulations, Title 27, Alcohol, Tobacco Products, and Firearms, paragraph 251.42a.  1990.

2.		PRINCIPLE

2.1		The measurement of excess pressure in a wine bottle has a direct correlation to the amount of CO2 in the wine.

3.		APPARATUS

3.1		Aphrometer.

3.1.1	Aphrometer Bouchon, Ligapal, Cormontreuil, France.

3.1.2	Modified Piercing Device, Zahm and Nagel Co., Inc., Buffalo, NY

3.2		Masterline Model 2095 Bath and Circulator, Forma Scientific, Marietta, OH

3.3		Branson Ultrasonic Bath, Model B-52, Branson Cleaning Equipment Co., Shelton, CT.

4.		CALIBRATION OF 					APHROMETER

The aphrometer was used as obtained from the supplier.  At this time, there is not an established gauge calibration procedure.

5.		EXPERIMENTAL PROCEDURE

5.1		General Procedures.

5.1.1.	Select the aphrometer.  Considering its ease of use, the Bouchon aphrometer is preferred if it fits over the stoppering device.

5.1.2.	Place the wine bottle to be tested in a water bath maintained at a temperature of 20 ± 0.1 °C.  Be sure that the water level of the bath is above the level of the wine in the bottle.  Allow it to equilibrate for one hour.

5.1.3	Remove the wine bottle from the bath.  Center the aphrometer on the stoppering device (cork or plastic).  Pierce the stoppering device as described below for each aphrometer and record the pressure.  Vent the pressure gauge to zero.  Reclose the system.

5.1.4	Place the wine bottle with attached aphrometer in an ultrasonic bath and sonicate for five minutes.

5.1.5	Replace the system in the thermostated water bath and allow it to equilibrate for thirty minutes.  Record the pressure a second time.  The second measurement is made to ensure that the excess pressure is due to carbon dioxide.  Remove the aphrometer.

5.1.6	For a wine to be considered sparkling, both pressure readings must be three bars or more.  It is recommended that a minimum of three bottles be tested.

5.2		Specific Procedures for the Aphrometer Bouchon.

5.2.1	This device (Figure 1) consists of a pressure gauge attached to a hand wheel that drives a hollow plunger equipped with a removable tip.  The gauge reads from 0 to 10 bars in one-tenth bar increments.

5.2.2	Secure the device to the wine bottle to be tested using the collar with four violin type screws.

5.2.3	Obtain a pressure reading by using the hand wheel to push the plunger and tip through the stoppering device.  Once through, the removable tip falls off and a reading of excess pressure is taken.  This procedure works smoothly and easily with both plastic and cork stoppering devices.

5.2.4	To prepare for the second pressure reading, the plunger is retracted with the hand wheel until it just comes free and then reinserted.  Both cork and plastic stoppers reseal without difficulty.

One drawback noted with this device is that the collar will accommodate only those stoppers with a diameter of 32 mm or less.

5.3		Specific Procedures for the Modified Piercing Device.

This device (Figure 2) consists of a base with two guides for the wings of a ventable pressure gauge attached to a hollow needle with four holes in it near its point.  The gauge reads 0-4 bars in one-tenth bar increments.

5.3.2	Hollow plastic champagne-type stopper.

5.3.2.1	To obtain a reading of excess pressure, the needle is easily forced through the stopper and holds its seal.

5.3.2.2	Pull the needle out after the reading has been taken.

5.3.3	Champagne cork or a regular cork

5.3.3.1	When the stoppering device is a real champagne cork or a regular cork, considerable force is necessary to pierce it.

5.3.3.2	Carefully measure the cork and mark a distance one-half inch from the bottom of the cork.

5.3.3.3	Using a drill bit with a slightly smaller diameter than that of the needle, mark this distance on the drill bit.

5.3.3.4	Carefully drill a hole into the cork to within one-half inch of the bottom of the cork.  Note:  Perform drilling before undertaking temperature equilibration of the bottle.

5.3.3.5	To obtain a reading of excess pressure, the needle of the aphrometer is forced through the remaining cork.  Record the pressure.

5.3.3.6	Vent the system in preparation for obtaining the second pressure reading.

5.3.3.7	After all pressure readings are taken, remove the needle from the cork. This sometimes requires the use of two persons—one to hold the bottle and one to pull out the needle.  Care should be taken for the safety of the involved personnel in both the piercing and removal parts of this procedure.  Heavy gloves are recommended as a minimum precaution.

5.3.3.8	Because of its hollow construction, the needle can be bent easily when it is being extracted.  The holes in the needle also tend to clog with cork particles.  This necessitates a manual cleaning with a stiff wire and compressed air before the needle can be reused.

5.3.3.9	The height of the guides precludes their use with bottles taller than thirty centimeters.

END