U.S. CUSTOMS AND BORDER PROTECTION

LABORATORY METHODS

CBPL METHOD 71-04

Gold Fineness in Articles of Jewelry

SAFETY STATEMENT

This CBPL Method cannot fully address safety issues that may arise from its use.  The analyst is responsible for assessing potential safety issues associated with a given method at its point of use.

Before using this method, the analyst will consider all general laboratory safety precautions.  In particular, the analyst will identify and implement suitable health and safety measures and will comply with all pertinent regulations.

METHOD UNCERTAINTY

The uncertainty of measurement for this method is specific to each laboratory.

0.		SCOPE AND FIELD OF 					APPLICATION

These procedures may be used in the analysis of articles of gold used for personal adornment for value determination for tariff purposes.  These articles, covered under the Harmonized Tariff System of the United States (HTSUS) numbers 7113.19.10-7113.19.50, 7114.19.00 and 7115.90.10, include small objects of personal adornment (gem-set or not) such as rings, bracelets, necklaces, brooches, ear-rings, neck chains, watch-chains, and other ornamental chains; fobs, pendants, tie-pins and clips, cuff-links, dress studs, buttons, etc; religious or other crosses; medals and insignia; hat ornaments (pins, buckles, rings etc.); ornaments for handbags; buckles and slides for belts, shoes, etc; hair-slides, tiaras, dress combs and similar hair ornaments.

If the gold content of the article is required to determine if it meets the standard of fineness (the gold content expressed as parts per thousand) required by the Gold Labeling Act of 1976 (public Law 94-450), the recommended method is CBPL Method 71-01, “Gold:  Quantitative Analysis by Fire Assay.”

1.		REFERENCES

1.1		Gold Labeling Act of 1976:  Public Law 94-450.

1.2		Furman, N.H. (Editor).  "Precipitation with Sulfurous Acid" in Scott’s Standard Methods of Chemical Analysis, Volume One.  D. Van Nostrand Co., Inc.  New York.  1962.  P 472.

2.		PRINCIPLE

The two procedures that are described may be used in the analysis of articles of gold used for personal adornment.  The procedure of choice depends upon the questions to be answered.  First is a non-destructive screening procedure using X-ray fluorescence.  The second is a gravimetric determination.

3.		APPARATUS, REGENTS, AND 			MATERIALS

3.1		Energy dispersive X-ray spectrometer system with a fundamental parameters program for XRF analysis.

3.2		Hydrochloric acid, reagent grade.

3.2.1	Hydrochloric acid, 1% (v/v) solution.

3.3		Nitric acid, reagent grade.

3.4		Aqua regia:  4 parts HCL, 1 part HNO3.

3.5		Sulfurous acid, saturated, reagent grade.

3.6		Filter paper, close-grained:  Whatman N. 44.

3.7		Standard laboratory equipment: including hot plates, steam bath, drying oven, filter funnels, Bunsen burners, graduated cylinders, litmus paper, beakers, Erlenmeyer flasks.

3.8		Analytical balance, ±0.0001 gram.

4.		SAMPLE PREPARATION

4.1		Analysis by XRF:  Clean the surface of any dust or dirt Remove any organic coating.

4.2		Analysis by sulfurous acid precipitation:  First clean the surface of any dust or dirt.

4.2.2	Cut a portion of the sample off using cutters or other appropriate means.  This portion should weigh approximately 1-2 grams.

4.2.3	Weigh each portion to ±0.0001 grams.

5.		PROCEDURE FOR ANALYSIS BY 		XRF

5.1		Set up XRF instrument as directed by manufacturer's directions.  The following are general guidelines:

a)  X-ray current is adjusted to give approximately 50% dead time.

b)  As a secondary target use gadolinium or other suitable target.

c)  Acquisition time is 100 seconds after correction for dead time.

5.2		Follow manufacturer's directions for operation of instrument and fundamental parameters program for XRF analysis.

5.3		Follow manufacturer's directions for calibration.

5.4		Choose a portion of the sample with a flat surface and the greatest area.  The surface presented to the X-ray beam should be homogeneous and not coated with organic material.

6.		PROCEDURE FOR 						PRECIPITATION WITH 					SULFUROUS ACID

6.1		Dissolve the sample in aqua regia.

6.2		After solution of the alloy or residue in aqua regia, any insoluble matter is filtered from the solution.

6.3		Carefully evaporate the solution on a steam bath.

6.4		Two additions, each of 5 mL of hydrochloric acid may be necessary to decomposed and volatilize the nitrates.

6.5		Dilute the solution to contain not more than 100 mg of gold per 100 mL.

6.6		Let the solution stand for several hours (See note 7.2).

6.7		Remove any insoluble matter by filtration.

6.8		Heat the clear solution to 80-90 °C.

6.9		Slowly add 20 mL of a saturated solution of sulfurous acid.

6.10		Heat the solution gently until the gold precipitate coagulates.

6.11		Add an additional 5 mL of the sulfurous acid solution and continue the heating for 10 minutes to insure complete precipitation of gold.

6.12		Filter the cooled solution a close-grained filter paper.

6.13		Wash the paper thoroughly with warm 1% hydrochloric acid and then to neutrality with water.

6.14		Dry the paper and ignite.

6.15		After cooling, weigh the precipitate of gold.

7.		NOTES ON THE PROCEDURE

7.1		Previous experience with the XRF method showed deviations of approximately ±1 karat or percent.

7.2		Care must be taken in filtering gold chloride not to expose the solution in contact with the filter paper to sunlight since gold is photosensitive and under such conditions will be reduced in the filter paper.

END